## DMM test results
Test & Measurement World
"Get your hands around a DMM," February 2003
Results from tests of handheld digital multimeters.
A recently calibrated HP (Agilent) 34401A bench DMM served as a “golden” meter for comparison of the handheld units. A Stanford Research DS335 function generator provided the 5.2-V and 9-V test signals, and a Tektronix TDS2022 oscilloscope supplied a view of the test waveforms.
Manufacturer | Model | Price (US$) | Number of counts | Accessories for probes | Temperature measurement | Lighted display/ visibility | Serial port | Bar graph | Dual display | Holster | Auto shutoff (min.) | DCV measurement (34401A) | DCV measurement (UUT) | Difference | Triangle ACV measurement (34401A) | Triangle ACV measurement (UUT) | Difference | AC line measurement 34401A | AC line measurement (UUT) | Difference
Group 1: $49.95
AW Sperry Instruments | DM-370A | 49.95 | 2000 | none | no | no | no | no | no | yes | no auto shutoff | 9.9088 | 9.89 | 0.0188 | 5.2078 | 5.15 | 0.0578 | 118.52 | 117.2 | 1.32
Group 2: $99 to $109
Extech Instruments | 22-316 | 99 | 4000 | probe tip covers | yes | yes/bright | no | no | no | yes | 15 | 9.909 | 9.89 | 0.019 | 5.2078 | 5.2 | 0.0078 | 118.46 | 118.4 | 0.06
Fluke | 110 | 99 | 6600 | none | no | no | no | yes | no | yes | 20 | 9.9088 | 9.92 | -0.0112 | 5.2078 | 5.208 | -0.0002 | 118.31 | 118.3 | 0.01
Greenlee Textron | DM-110 | 99 | 4000 | alligator clip (black only) | no | no | no | yes | yes | yes | 30 | 9.9086 | 9.92 | -0.0114 | 5.2078 | 5.24 | -0.0322 | 118.81 | 118.9 | -0.09
Morrow Wave | DMM-531T | 99.95 | 4000 | alligator clips | no | no | no | yes | no | yes | 30 | 9.9088 | 9.92 | -0.0112 | 5.2078 | 5.26 | -0.0522 | 118.46 | 119.3 | -0.84
Meterman | 34XR | 109.95 | 4000 | alligator clips and tip covers | yes | yes/bright | no | yes | no | yes | 30 | 9.9088 | 9.91 | -0.0012 | 5.2078 | 5.23 | -0.0222 | 118.23 | 117.9 | 0.33
Group 3: $129 to $145.95
Fluke | 111 | 129 | 6600 | none | no | no | no | yes | no | yes | 20 | 9.9087 | 9.92 | -0.0113 | 5.2078 | 5.209 | -0.0012 | 118.4 | 118.3 | 0.1
Extech Instruments | MT330 | 129 | 4000 | alligator clips | yes | yes/dim | yes (cable and software optional) | no | no | yes | 30 | 9.9087 | 9.92 | -0.0113 | 5.2078 | 5.26 | -0.0522 | 118.86 | 119.5 | -0.64
Brighton Electronics | Summit 610 | 129.95 | 4000 | alligator clips | no | no | no | yes | no | yes | 30 | 9.9086 | 9.92 | -0.0114 | 5.2078 | 5.25 | -0.0422 | 118.75 | 119 | -0.25
Global Specialties | Pro-80 | 135.95 | 80000 | probe tip covers | yes (probe not included) | yes/dim | non-isolated | yes |  | yes | 30 | 9.9088 | 9.905 | 0.0038 | 5.2079 | 5.2105 | -0.0026 | 118.68 | 119.25 | -0.57
Brighton Electronics | Summit 620 | 145.95 | 4000 | alligator clips | yes | no | no | yes | no | yes | 45 | 9.9087 | 9.92 | -0.0113 | 5.2078 | 5.23 | -0.0222 | 118.63 | 119 | -0.37
Group 4: $169 to $199
Protek | 506 | 169 | 4000 | alligator clips | yes | yes/bright | non-isolated |  | yes | no | 30 | 9.9087 | 9.91 | -0.0013 | 5.2078 | 5.26 | -0.0522 | 118.53 | 119.4 | -0.87
Global Specialties | Pro-60 | 179.95 | 8000 | none | yes | yes/bright | isolated (cable and software included) | yes | yes | yes | 10 | 9.9076 | 9.87 | 0.0376 | 5.2095 | 5.207 | 0.0025 | 118.53 | 118.5 | 0.03
B&K Precision | 391A | 187 | 20000 | none | no | no | no | no | no | yes | 45 | 9.9087 | 9.913 | -0.0043 | 5.2078 | 5.254 | -0.0462 | 118.55 | 119 | -0.45
B&K Precision | 2880A | 199 | 40000 | none | no | no | yes (cable and software optional) | no | yes | yes | 15 | 9.9091 | 9.91 | -0.0009 | 5.2078 | 5.25 | -0.0422 | 118.46 | 119 | -0.54
Group 5: $219 to $249
AEMC Instruments | CA 5240 | 219 | 40000 | none | no | yes/bright | no | yes | no | yes | 30 | 9.9085 | 9.902 | 0.0065 | 5.2077 | 5.338 | -0.1303 | 118.33 | 119.29 | -0.96
Omega Engineering | HHM290 | 245 | 40000 | none | yes (2 probes) | yes/bright | no | yes | yes | yes | 20 | 9.9089 | 9.906 | 0.0029 | 5.2078 | 5.203 | 0.0048 | 118.28 | 118.28 | 0
Protek | 608 | 249 | 50000 | alligator clips | yes (probe optional) | yes | isolated (cable and SW optional) | yes | yes | yes | 15 | 9.9091 | 9.9 | 0.0091 | 5.2078 | 5.229 | -0.0212 | 118.47 | 118.6 | -0.13